## Sheet1
Technical Datasheet RFQ - 04/2026
No | Technical Specification | Supplier offered specification | Compliant (yes/No) | Remarks / Deviation Details | Supporting Document
1 | Major Surgical Instrument Set
1. General Description
Major surgical instrument set suitable for general surgical procedures in hospitals and operating rooms.
2. Material
| Instruments made of high-quality stainless steel |  |  |  |  | uuuu
Rust-resistant and suitable for repeated sterilization
Smooth finish with durable construction
3. Set Components and Sizes
Dressing forceps: approximately 14–16 cm
Tissue forceps: approximately 14–18 cm
Artery forceps (straight and curved): approximately 14–20 cm
Kocher forceps: approximately 14–18 cm
Allis tissue forceps: approximately 14–16 cm
Needle holder: approximately 16–20 cm
Mayo scissors (straight and curved): approximately 14–18 cm
Metzenbaum scissors: approximately 14–18 cm
Sponge holding forceps: approximately 18–24 cm
Backhaus towel forceps: approximately 9–13 cm
Scalpel handle: No. 3 and/or No. 4
Retractors: small and medium sizes
Suction tube/cannula (if applicable)
Kidney tray or instrument tray: standard size
Equivalent standard surgical instruments and sizes are acceptable
4. Sterilization Compatibility
Suitable for autoclave sterilization
Resistant to corrosion and repeated cleaning/disinfection processes
5. Storage and Packing
Supplied in sterilization tray, pouch, or carrying case
Instruments properly packed and protected during transport
6. Performance Requirements
Smooth movement and proper alignment of instruments
Comfortable handling and reliable clinical performance
Suitable for routine major surgical procedures
7. Warranty
1-year warranty against manufacturing defects
8. Spare Parts / Replacement
Availability of replacement instruments when required
Supplier shall support replacement of damaged or worn components
9. Quality Standards
ISO 13485 compliant or equivalent
CE Mark and/or FDA approval preferred
Compliance with international surgical instrument standards
10. Technical Documents
Product catalogue and detailed list of components
Material specification (stainless steel grade)
Quality certificates (ISO, CE, FDA or equivalent)
Warranty statement
2 | MINOR SURGICAL SET
Technical Specification
1. General Description
The Basic Minor Surgical Set is a simplified of reusable surgical instruments intended for use in primary healthcare facilities, outpatient clinics, dressing rooms, and emergency first-aid procedures.
The set is designed to meet basic clinical requirements such as minor wound care, simple suturing, dressing, and basic tissue handling procedures.
All instruments shall be manufactured from medical-grade stainless steel suitable for repeated cleaning and sterilization.
2. Technical Specifications
General Technical Requirements
Instruments shall be manufactured from medical stainless steel (AISI 420 or equivalent).
Corrosion-resistant and suitable for repeated use.
Compatible with autoclave sterilization at 134°C.
Satin or matt finish to reduce glare under clinical lighting.
Smooth surfaces free from manufacturing defects or sharp edges.
Basic ergonomic design suitable for routine clinical use.
Simple and reliable locking mechanisms where applicable.
Scissors shall provide adequate cutting performance for soft tissue and dressing materials.
Forceps tips shall ensure proper grip for routine procedures.
Instruments shall withstand repeated sterilization cycles without deterioration.
Equivalent products from internationally recognized manufacturers shall be acceptable.
3. Set Components
Scalpel Handle No. 3 — Qty: 1
Scalpel Handle No. 4 — Qty: 1
Mayo Scissors Straight, length 14–15 cm — Qty: 1
Mosquito Hemostatic Forceps Straight, length 12–13 cm — Qty: 2
Mosquito Hemostatic Forceps Curved, length 12–13 cm — Qty: 2
Needle Holder Mayo Hegar, length 14–16 cm — Qty: 1
Dressing Forceps Standard, length 14–16 cm — Qty: 1
Tissue Forceps Standard, length 14–16 cm — Qty: 1
Backhaus Towel Clamp, length 11–13 cm — Qty: 2
Stainless Steel Kidney Dish, capacity approximately 500–700 ml — Qty: 1
Stainless Steel Gallipot, capacity approximately 120–200 ml — Qty: 1
Stainless Steel Instrument Tray, size approximately 25 × 20 × 5 cm — Qty: 1
4. Warranty Requirements
Warranty period shall be a minimum of twenty-four (24) months from the date of delivery and acceptance.
Warranty shall cover manufacturing defects and material failures under normal use.
Supplier shall replace defective instruments during the warranty period without additional cost.
After-sales support shall be available throughout the warranty period.
5. Quality Certificates
2 | The supplied instruments shall comply with international medical device standards and include valid copies of:
ISO 13485 – Medical Devices Quality Management System
CE Marking or equivalent certification
ISO 9001 (preferred)
Declaration of Conformity
Certificate of Origin
All certificates shall be valid and issued by recognized certification bodies.
6. Required Documents
The supplier shall provide the following documents:
Manufacturer catalogue or technical datasheet
Certificate of Origin
Packing List
User and maintenance instructions
Warranty certificate
Sterilization compatibility statement
Delivery note and inspection report
7. Inspection and Acceptance
Acceptance shall be based on:
Visual inspection of instruments
Verification of quantities and specifications
Functional performance check
Review of submitted documents and certificates
Compliance with all technical requirements
3 | Orthopedic Surgical Instrument Set
1. General Description
Orthopedic surgical instrument set suitable for general orthopedic procedures in hospitals and operating rooms.
2. Material
Instruments made of high-quality stainless steel
Rust-resistant and suitable for repeated sterilization
Durable construction with smooth finish
3. Set Components and Sizes
Bone holding forceps: approximately 18–24 cm
Reduction forceps: approximately 16–22 cm
Periosteal elevator: approximately 14–18 cm
Bone curettes: assorted sizes
Osteotome/chisel: assorted sizes
Bone mallet: medium size
Gigli saw handle with wire: standard size
Rongeur forceps: approximately 16–20 cm
Bone cutter: approximately 18–22 cm
Orthopedic retractors: assorted sizes
Drill chuck/key (manual type if applicable)
Screwdriver set for orthopedic use
Depth gauge: standard size
Plate holding forceps: approximately 18–20 cm
Stainless steel tray or instrument case: standard size
Equivalent standard orthopedic instruments and sizes are acceptable
4. Sterilization Compatibility
Suitable for autoclave sterilization
Resistant to corrosion and repeated cleaning/disinfection processes
5. Storage and Packing
Supplied in sterilization tray, pouch, or carrying case
Instruments properly packed and protected during transport
6. Performance Requirements
Smooth movement and proper alignment of instruments
Strong and durable for orthopedic surgical use
Suitable for routine orthopedic procedures
7. Warranty
1-year warranty against manufacturing defects
8. Spare Parts / Replacement
Availability of replacement instruments when required
Supplier shall support replacement of damaged or worn components
9. Quality Standards
ISO 13485 compliant or equivalent
CE Mark and/or FDA approval preferred
Compliance with international orthopedic surgical instrument standards
10. Technical Documents
Product catalogue and detailed list of components
3 | Material specification (stainless steel grade)
Quality certificates (ISO, CE, FDA or equivalent)
Warranty statement
4 | OBSTETRICS & GYNECOLOGY (OB/GYN) SURGICAL SET
Technical Specification
1. General Description
The Obstetrics and Gynecology (OB/GYN) Surgical Set is a comprehensive collection of reusable surgical instruments intended for use in obstetric and gynecological procedures within operating rooms, maternity wards, delivery rooms, and emergency surgical settings.
The set is designed to support a wide range of procedures including vaginal examination, assisted delivery procedures, episiotomy repair, uterine exploration, dilation and curettage, and minor gynecological surgical interventions.
All instruments shall be manufactured from high-quality medical-grade stainless steel suitable for repeated sterilization and long-term clinical use.
2. Technical Specifications
General Technical Requirements
Instruments shall be manufactured from medical stainless steel (AISI 420 / AISI 304 or equivalent).
Corrosion-resistant and suitable for repeated sterilization.
Compatible with autoclave sterilization at 134°C.
Satin or matt finish to reduce glare under clinical lighting.
Smooth surfaces free from manufacturing defects, burrs, or sharp edges.
Ergonomic design for safe and precise handling.
All joints, ratchets, and locking mechanisms shall operate smoothly and reliably.
Forceps and clamps shall ensure accurate alignment and secure grip.
Retractors shall provide adequate exposure of the operative field.
Cutting instruments shall maintain sharp and durable performance.
Instruments shall withstand repeated use and sterilization without degradation.
Equivalent products from internationally recognized manufacturers shall be acceptable.
3. Set Components
Sims Vaginal Speculum, medium size — Qty: 1
Cusco Vaginal Speculum, medium size — Qty: 1
Sponge Holding Forceps, length 20–24 cm — Qty: 2
Ovum Forceps, length 24–26 cm — Qty: 2
Allis Tissue Forceps (4×5 teeth), length 14–16 cm — Qty: 2
Foerster Uterine Forceps, length 24–26 cm — Qty: 2
Sponge Forceps Straight/Curved, length 20–24 cm — Qty: 2
Needle Holder Mayo Hegar, length 16–20 cm — Qty: 2
Episiotomy Scissors, curved, length 14–16 cm — Qty: 1
Mayo Scissors Curved, length 14–17 cm — Qty: 1
Uterine Curette (sharp), sizes assorted — Qty: 2
Uterine Curette (blunt), sizes assorted — Qty: 2
Uterine Dilators (Hegar type), set of graduated sizes — Qty: 1 set
Sims Uterine Sound, length 25–30 cm — Qty: 1
Stainless Steel Kidney Dish — Qty: 2
Stainless Steel Gallipot — Qty: 4
Stainless Steel Instrument Tray — Qty: 1
The complete set shall be supplied in a sterilizable stainless steel container or sterilization tray suitable for safe storage, transport, and autoclave processing.
4. Warranty Requirements
Warranty period shall be a minimum of twenty-four (24) months from the date of delivery and acceptance.
Warranty shall cover manufacturing defects, material defects, and functional failures under normal use conditions.
Supplier shall replace defective instruments during the warranty period without additional cost.
Full after-sales service and technical support shall be provided throughout the warranty period.
Availability of maintenance support and spare parts shall be ensured.
5. Quality Certificates
The supplied instruments shall comply with international medical device standards and include valid copies of:
ISO 13485 – Medical Devices Quality Management System
CE Marking / CE Certificate
FDA Registration or Approval (if applicable)
ISO 9001 Certification (preferred)
Declaration of Conformity
Certificate of Origin
All certificates must be valid and issued by recognized international certification bodies.
6. Required Documents
The supplier shall provide the following documents:
Manufacturer catalogue and technical datasheets
Certificate of Origin
Packing List
User and maintenance manual
Warranty certificate
Sterilization compatibility statement
Manufacturer authorization letter (if applicable)
Delivery note and inspection report
7. Inspection and Acceptance
Acceptance shall be based on:
Visual inspection of instruments
Verification of quantities and specifications
Functional performance testing
Review of submitted documents and certificates
Compliance with all technical requirements
5 | ENT Surgical Instrument Set
1. General Description
ENT (Ear, Nose & Throat) surgical instrument set suitable for routine ENT procedures in hospitals, clinics, and operating rooms.
4 | 2. Material
Instruments made of high-quality stainless steel
Rust-resistant and suitable for repeated sterilization
Smooth finish with durable construction
3. Set Components and Sizes
Nasal speculum: assorted sizes
Ear speculum: assorted sizes
Tongue depressor: standard size
Tuning fork set: standard medical type
Dressing forceps: approximately 14–16 cm
Tilley forceps: approximately 15–18 cm
Artery forceps (straight and curved): approximately 14–18 cm
ENT scissors: approximately 14–18 cm
Needle holder: approximately 14–16 cm
Suction tube/cannula: assorted sizes
Retractors: small sizes
Scalpel handle: No. 3 and/or No. 4
Kidney tray or instrument tray: standard size
Equivalent standard ENT instruments and sizes are acceptable
4. Sterilization Compatibility
Suitable for autoclave sterilization
Resistant to corrosion and repeated cleaning/disinfection processes
5. Storage and Packing
Supplied in sterilization tray, pouch, or carrying case
Instruments properly packed and protected during transport
6. Performance Requirements
Smooth movement and proper alignment of instruments
Comfortable handling and reliable clinical performance
Suitable for routine ENT procedures
7. Warranty
1-year warranty against manufacturing defects
8. Spare Parts / Replacement
Availability of replacement instruments when required
Supplier shall support replacement of damaged or worn components
9. Quality Standards
ISO 13485 compliant or equivalent
CE Mark and/or FDA approval preferred
Compliance with international ENT surgical instrument standards
10. Technical Documents
Product catalogue and detailed list of components
Material specification (stainless steel grade)
Quality certificates (ISO, CE, FDA or equivalent)
Warranty statement
6 | Urology Surgical Instrument Set
1. General Description
Urology surgical instrument set suitable for routine urological procedures in hospitals and operating rooms.
2. Material
Instruments made of high-quality stainless steel
Rust-resistant and suitable for repeated sterilization
Smooth finish with durable construction
3. Set Components and Sizes
Urethral dilators: assorted sizes
Metal urinary catheters: assorted sizes
Sponge holding forceps: approximately 18–24 cm
Artery forceps (straight and curved): approximately 14–18 cm
Tissue forceps: approximately 14–16 cm
Needle holder: approximately 16–18 cm
Surgical scissors (straight and curved): approximately 14–18 cm
Retractors: small and medium sizes
Suction tube/cannula: assorted sizes
Kidney stone forceps (if applicable): standard size
Scalpel handle: No. 3 and/or No. 4
Kidney tray or instrument tray: standard size
Equivalent standard urology instruments and sizes are acceptable
4. Sterilization Compatibility
Suitable for autoclave sterilization
Resistant to corrosion and repeated cleaning/disinfection processes
5. Storage and Packing
Supplied in sterilization tray, pouch, or carrying case
Instruments properly packed and protected during transport
6. Performance Requirements
Smooth movement and proper alignment of instruments
Comfortable handling and reliable clinical performance
Suitable for routine urology procedures
7. Warranty
1-year warranty against manufacturing defects
8. Spare Parts / Replacement
Availability of replacement instruments when required
Supplier shall support replacement of damaged or worn components
9. Quality Standards
ISO 13485 compliant or equivalent
CE Mark and/or FDA approval preferred
Compliance with international urology surgical instrument standards
10. Technical Documents
Product catalogue and detailed list of components
Material specification (stainless steel grade)
Quality certificates (ISO, CE, FDA or equivalent)
Warranty statement
7 | Defibrillator
1. General Description
Portable or semi-automatic defibrillator used for emergency resuscitation of patients with cardiac arrest or life-threatening arrhythmias.
2. Type
Manual and/or AED (Automated External Defibrillator) type
Suitable for hospital emergency rooms, ambulances, and clinics
3. Energy Delivery
Biphasic waveform technology
Energy range: approximately 1–200/360 Joules (depending on model)
Adjustable energy levels for adult and pediatric use
4. Display and Controls
Clear LCD or LED display showing ECG and device status
Simple control panel for emergency operation
Audio and visual prompts (for AED models)
5. Monitoring Functions
ECG monitoring capability
Heart rate display
Battery and system status indicators
6. Power Supply
Rechargeable internal battery
AC power supply: 220–240V, 50/60 Hz
Battery backup for emergency operation
7. Safety Features
Automatic self-check system
Shock delivery safety confirmation
Patient isolation and leakage protection
Audible alarms for faults or incorrect operation
8. Accessories
Defibrillation paddles or pads (adult & pediatric)
ECG cables and electrodes
Rechargeable battery and charger
Carrying case (if portable type)
9. Environmental Requirements
Operating temperature: 10°C – 40°C
Humidity: up to 85% non-condensing
Suitable for standard hospital environments and emergency transport use
10. Warranty
2 years full warranty covering manufacturing and electronic faults
11. Spare Parts
Availability of spare parts for 10 years minimum
Includes pads, batteries, cables, and internal components
12. Quality Standards
ISO 13485 (mandatory)
ISO 9001 preferred
CE Mark and/or FDA approval required or equivalent certification
Compliance with international medical electrical safety standards
13. Technical Documents
Product catalogue and datasheet
Quality certificates (ISO 13485, CE, FDA or equivalent)
Manufacturer authorization (if applicable)
Warranty and spare parts confirmation
List of similar installations (if available)
8 | Mobile Operating Lamp
1. General Description
Mobile LED operating lamp designed to provide high-intensity, shadow-free illumination for minor surgical procedures and clinical examinations.
2. Type
Mobile / portable floor-standing surgical light
Suitable for operating rooms, emergency rooms, and clinics
3. Light Source
LED technology
High illumination intensity with long lifespan
Shadow-free lighting system
4. Illumination
Adjustable light intensity
High brightness suitable for surgical procedures
Color temperature suitable for medical use (white daylight range)
5. Mobility and Structure
Mounted on stable mobile base with wheels
Easy movement and positioning
Lockable wheels for stability during use
Adjustable arm for flexible positioning
6. Power Supply
AC power supply: 220–240V, 50/60 Hz
Low power consumption
Backup option (if available) via battery system (optional)
7. Control System
Manual brightness adjustment
On/off control panel
Easy-to-use interface for clinical environments
8. Safety and Performance
Low heat emission to ensure patient comfort
Stable illumination without flickering
Designed for continuous medical use
9. Environmental Requirements
Operating temperature: 10°C – 40°C
Humidity: up to 85% non-condensing
Suitable for standard hospital environments
10. Warranty
2 years full warranty covering manufacturing and electrical faults
11. Spare Parts
Spare parts availability for 10 years minimum
Includes LED modules, power supply units, and control components
12. Quality Standards
ISO 13485 (mandatory)
ISO 9001 preferred
CE Mark and/or FDA approval required or equivalent certification
Compliance with medical electrical safety standards
13. Technical Documents
Product catalogue and technical datasheet
Quality certificates (ISO, CE, FDA or equivalents)
Manufacturer authorization (if applicable)
Warranty statement
Spare parts availability confirmation
List of similar installations (if available)
9 | Portable Ultrasound Machine (1 Convex Probe, 2 Probe Sockets)
1. General Description
Portable ultrasound diagnostic system suitable for abdominal, obstetric, and general imaging in hospitals, clinics, and emergency settings.
2. System Type
Portable ultrasound unit (laptop or trolley type)
Supports at least two active probe sockets
Suitable for bedside and outpatient use
3. Probes
Supplied with 1 Convex probe (standard abdominal/OB use)
System must have minimum 2 active probe ports (sockets)
Additional probes should be optional or expandable
4. Imaging Modes
B-mode (2D imaging)
M-mode (preferred)
Color Doppler (preferred)
PW / Power Doppler (optional but preferred)
5. Display
High-resolution medical LCD/LED screen
Real-time imaging display
Adjustable brightness and viewing angle
6. Performance Functions
Freeze and store images
Measurement and calculation tools
Image zoom and review function
Fast boot and real-time response
7. Power Supply
AC power: 220–240V, 50/60 Hz
Built-in rechargeable battery for portable use
Battery backup for emergency operation
8. Data Storage & Connectivity
Internal storage for images and patient data
USB export function
DICOM compatibility (preferred)
9. Safety & Environment
Suitable for continuous clinical use
Operating temperature: 10°C – 40°C
Humidity: up to 85% non-condensing
Low power consumption and stable performance
10. Warranty
2 years full warranty covering manufacturing and electronic defects
11. Spare Parts
Spare parts availability for 10 years minimum
Includes probes, cables, power unit, and internal components
12. Quality Standards
ISO 13485 (mandatory)
ISO 9001 preferred
CE Mark and/or FDA approval required or equivalent certification
Compliance with international medical imaging standards
13. Technical Documents
Product catalogue and full technical datasheet
Probe specifications and system configuration
Quality certificates (ISO, CE, FDA or equivalent)
Manufacturer authorization (if applicable)
Warranty statement
Spare parts availability confirmation
List of previous installations (if available)
10 | Electrosurgical Diathermy Unit with Trolley
1. General Description
Electrosurgical diathermy unit suitable for cutting and coagulation in general surgical procedures, supplied with a mobile trolley for easy mobility and clinical use.
2. System Type
Monopolar and bipolar electrosurgical unit (preferred)
Supplied with dedicated mobile trolley
Suitable for operating rooms and minor surgical procedures
3. Output Functions
Pure cut mode
Coagulation mode
Blended cut/coagulation mode
Bipolar coagulation mode (preferred)
4. Power Output
Adjustable output power for different surgical needs
Stable energy delivery with automatic power control
Suitable for soft tissue and general surgical applications
5. Control System
Digital or manual control panel
Clear display of power settings and mode selection
Easy and safe operation for clinical staff
6. Safety Features
Patient return electrode monitoring system
Alarm system for faulty connection or safety issues
Overload and overheating protection
Automatic self-check system on startup
7. Trolley
Mobile medical-grade trolley included
Strong metal frame with stable base
Equipped with wheels (lockable) for mobility and safety
Space for accessories and foot switch
8. Accessories
Patient plate (return electrode)
Foot switch pedal
Electrosurgical handpiece with electrodes
Connection cables and basic accessories
9. Power Supply
AC power: 220–240V, 50/60 Hz
Stable operation under hospital electrical conditions
Built-in protection against voltage fluctuations
10. Safety & Environmental Requirements
Suitable for continuous clinical use
Operating temperature: 10°C – 40°C
Humidity: up to 85% non-condensing
Designed for operating room safety standards
11. Warranty
2 years full warranty covering manufacturing and electrical faults
12. Spare Parts
Spare parts availability for 10 years minimum
Includes handpiece components, cables, and electronic modules
13. Quality Standards
ISO 13485 (mandatory)
ISO 9001 preferred
CE Mark and/or FDA approval required or equivalent certification
Compliance with international electrosurgical safety standards
14. Technical Documents
Product catalogue and technical datasheet
Quality certificates (ISO, CE, FDA or equivalent)
Manufacturer authorization (if applicable)
Warranty statement
Spare parts availability confirmation
List of previous installations (if available)
11 | Hydraulic Surgical Operating Table (Non-Electrical)
1. Description of Function
The operating table is designed to safely position and support patients during surgical procedures. It shall provide stable and precise mechanical adjustment of patient position using a hydraulic system, enabling optimal access for surgical teams and ensuring patient safety and comfort.
2. Technical Specifications
2.1 General Construction
The table frame shall be constructed from medical-grade stainless steel (AISI 304 or AISI 316).
The structure shall be corrosion-resistant and suitable for repeated hospital-grade disinfection.
The base shall be heavy-duty to ensure maximum stability during all surgical procedures.
2.2 Hydraulic System
The table shall operate using a fully mechanical hydraulic system (non-electrical).
Height adjustment shall be controlled via foot-operated hydraulic pump or hydraulic cylinder system.
The hydraulic system shall be a closed-loop, leak-proof design.
The table shall maintain its position securely without drift or unintended movement.
2.3 Table Top
The mattress shall consist of high-density medical-grade foam (40–80 mm thickness).
Cover material shall be antibacterial PU or PVC, resistant to fluids and disinfectants.
The surface shall be removable or easy to clean and sterilize.
2.4 Dimensions
Overall length: approx. 180–210 cm
Width: approx. 50–60 cm
Height adjustment range: approx. 70–100 cm
9 | 2.5 Load Capacity
Minimum safe working load: 200–250 kg
3. Required Movements
The operating table shall provide the following mechanical movements:
Height adjustment (up/down)
Trendelenburg position: up to 25°–30°
Reverse Trendelenburg: up to 25°–30°
Lateral tilt (left/right): up to 20°
Backrest adjustment: up to 70°
Leg section adjustment: up to 90°
All positions shall include a secure locking mechanism
4. Safety Requirements
Mechanical locking system to prevent unintended movement
Anti-slip base for maximum stability
High structural strength to prevent tipping under maximum load
Rounded edges to reduce injury risk
Non-flammable and hospital-safe materials
Designed for infection control and continuous disinfection
5. Standard Accessories
The table shall be supplied with at least the following accessories:
Adjustable headrest support
Arm boards (removable and adjustable)
Leg supports
Patient securing straps
Side rails (standard medical size)
High-density medical mattress
Anesthesia screen frame/bar
6. Installation & Commissioning
Full installation at the hospital site by qualified technicians
Verification of hydraulic system functionality
Testing of all movements and locking positions
Load testing under simulated operating conditions
Final acceptance upon successful commissioning tests
7. Training
On-site training for medical and technical staff
Operation of all hydraulic functions and positioning mechanisms
Emergency handling and safety procedures
Basic maintenance and daily inspection procedures
8. Warranty & Maintenance
Minimum warranty period: 24 months
Coverage includes hydraulic system and structural frame defects
Spare parts availability for a minimum of 10 years
Preventive maintenance schedule to be provided by the supplier
Response time for service support shall not exceed 24–48 hours from the time of notification
9. Documentation
The supplier shall provide complete documentation including:
User manual (operation and safety instructions)
Maintenance manual (preventive and corrective maintenance)
Technical drawings and hydraulic schematics
Load test and quality inspection reports
Spare parts list
Warranty certificates
10. Quality Standards & Certifications
The offered equipment shall comply with internationally recognized medical device standards:
ISO 13485 – Medical Device Quality Management System
ISO 14971 – Risk Management for Medical Devices
CE Marking or equivalent international regulatory approval
Certification of medical-grade stainless steel (AISI 304/316)
Certification confirming antibacterial, non-toxic PU/PVC materials
| Factory quality certification and audit report (if required) |  | Infant Incubator
Load testing and hydraulic system performance test reports
12 | Laryngoscope Set – Adult/Child
1. General Requirements
1. The set shall include both adult and pediatric laryngoscopes with appropriate blades for each age group.
2. The equipment shall be brand new, latest model (manufactured 2025 or later).
3. All components shall be autoclavable and suitable for repeated sterilization cycles without degradation.
4. The manufacturer shall have proven international experience in medical airway devices.
2. Regulatory & Standards Compliance
The equipment must comply with:
ISO 13485
ISO 10993 (biocompatibility)
EU MDR 2017/745 (CE Marking under MDR) or US FDA clearance
IEC 60601-1 (for battery-powered laryngoscopes)
All certificates and declarations must be submitted with the technical offer.
3. Laryngoscope Components
3.1 Handles
Material: Stainless steel or high-grade anodized aluminum
Ergonomic design for easy grip
Autoclavable
Compatible with both standard and pediatric blades
Power source: Rechargeable battery or disposable batteries (with at least 2 hours continuous illumination)
Switch: Reliable on/off toggle or push-button
Non-slip finish for safe handling
3.2 Blades
Adult Blades
Sizes: #3, #4 (Macintosh curved), #3, #4 (Miller straight)
Material: Stainless steel or high-strength autoclavable alloy
Light source: Fiber optic or LED illumination
Smooth, atraumatic edges
Pediatric Blades
Sizes: #0, #1, #2 (Macintosh and Miller options)
Material: Stainless steel or autoclavable alloy
Fiber optic or LED illumination
Smooth, atraumatic edges
3.3 Illumination
High-intensity LED light (minimum 50 lumens)
Consistent brightness throughout battery life
Backup light option in case of failure
3.4 Accessories
Rechargeable battery charger and spare batteries
Blade storage box or tray with separation for pediatric/adult blades
Protective blade covers or sheaths
Cleaning and sterilization instructions included
4. Operating & Storage Conditions
Operating temperature: 15°C – 35°C
Storage temperature: -5°C – 40°C
Relative humidity: 10–95% non-condensing
5. Documentation (Mandatory)
The supplier shall provide:
1. ISO 13485 certificate
2. CE MDR certificate or US FDA clearance
3. Declaration of Conformity
4. User (Operating) Manual – English
5. Service / Maintenance Manual – English
6. Cleaning, disinfection, and sterilization instructions
7. Spare parts list with part numbers
8. Warranty certificate
6. Warranty & After-Sales Service
Comprehensive warranty: Minimum 2 years
Preventive maintenance support
Spare parts availability: minimum 10 years
Service response time ≤ 72 hours
7. Acceptance Criteria
Final acceptance shall be based on:
Visual inspection for defects
Blade fit and handle compatibility
Illumination test
Battery performance check
Verification of submitted documentation
13 | Dental Machine
Description of Function
A Dental chair for Dental treatment
Operational Requirements
It shall operate on AC power supply.
Dental Unit with complete attachments and accessories.
Patient chair, 1 unit
Working stool, 1 unit
A built-in light cure unit
A built-in ultrasonic scaler with one each of pointed and flat scaler tips
One air compressor at least 1 horse power
Bidder shall indicate brand and model information here and provide technical data document for major components specified above.
Technical Specifications
Main unit standard configuration as follow:
Patient chair:
Electric chair movement and deluxe double articulated headrest.
With one left armrest as standard.
3 pre-set chair positions: start, treatment and rinsing position.
Electrical patient chair loading capacity: not less than 135kg.
Backrest movement range 105°-175°.
The lowest position of the patient chair from the ground shall not be less then 380mm.
The highest position of the patient chair from the ground shall not be less than 780mm.
Chair movement is controllable by the 4 way foot control at the chair base without touch panel.
No cables on the floor, hygienic and clean.
One main switch to control air, water and power.
The chair position is locked while an instrument is working.
With chair-backrest safety system, backrest and seat movement can be stopped once it meets obstacle.
Dentist element:
Dentist element with whip arm system.
Height of dentist element is adjustable.
1 X-ray film viewer (12V, 2000cd/m2).
1 silicon mat for the dentist element which can be sterilized.
1 three way syringe.
3 ISO 4-hole/Midwest hand piece hoses.
1 air pressure meter.
Assistant element:
1 three way syringe.
1 strong suction hose.
1 saliva ejector.
With suction filter system.
Water unit:
The cuspidor can be swivelled and removable for easy cleaning.
Cup filler and bowl rinsing systems shall prevent over filling of cup and prolong rinsing of bowl. Preferably programmable.
With automatic water heating system (24V)
With water venturi and air water separator system
Fresh water bottle, at least 1.5L
Operating light:
Colour temperature: 3800-4500K.
12V, 50 Watt Halogen bulb or Better.
Dental light intensity: min 25000lux with intensity dimming function.
Working stool, 1 units.
Mobile on 5 castors.
Height of seat and backrest is adjustable.
Backrest angle is adjustable and lockable.
Come with a built-in light cure unit.
Come with one suction unit
Come with a built-in ultrasonic scaler with one each of pointed and flat scaler tips.
Come with one air compressor at least 1 horse power or capacity sufficient to supply to the Dental chair specified above, whichever higher. Bidder shall indicate capacity of the unit offered here. Air compressor must be oil less & noise less
Accessories, Spare Parts and Consumables
All standard accessories, consumables and parts required to operate the equipment.
Power supply
Power supply: 220 – 240 VAC, 50Hz fitted with appropriate plug. The power cable must be at least 3 meters in length.
Standards & Safety Requirements
Must submit ISO13485:2003/AC:2007 for Medical Devices AND
CE (93/42 EEC Directives) or USFDA approved product certificate.
Electrical safety conforms to standards for electrical safety IEC-60601.
User Training
The Supplier shall conduct user training for this equipment to enable operators to use the equipment properly. The training shall include the use of all operational functions of the equipment, as well as routine checks and maintenance expected by users.
Warranty
The warranty period for this item shall be two years
Installation and Commissioning
The bidder must be installed and commissioned equipment
Documentation
User (Operating) manual in English
Service (Technical / Maintenance) manual in English
Country Of Origin
Manufacture / Brand
Model / Type
Delivery time
Total Price (USD)
14 | Digital X-ray Machine
Description of Function
A general purpose digital X-ray machine for routine X-ray examinations at healthcare facilities.
Operational Requirements
It shall be suitable to be used for adult and pediatric patients in general Radiography examination.
System Configuration
X-ray Generator,1 unit
X-Ray tube & tube support system, 1 unit
Radiographic patient table, fixed height1 unit
Floor mounted bucky stand, 1 unit
Technical Specifications
Should be floor to floor mounted
X-ray Generator:
Microprocessor based, high frequency inverter generator, the generator shall have at least 100kHz.
Generator Output: not less than 30 Kw
Radiographic voltage: 40 kV to 140kV, in 1Kv step.
Radiographic current approx., 10 to 500 mA
mAs range: 0.5 mAs to 650 mAs.
Anatomical Programmable Radiographic mode shall be available.
Manual & automatic exposure control and automatic brightness control shall be available.
Shall come with overload protection device.
X-Ray Tube:
Maximum tube output shall match with the generator output
X-ray tube rotating: +/-90°.
Anode rotating speed: 2700 rpm or more.
Large focus not more than 1.2 mm.
Small focus not more than 0.6 mm.
Maximum tube voltage 140 KV.
Max housing heat dissipation rate: 250HU/s
Anode heat storage capacity: min 140 kHU (100kJ);
Max. heat storage capacity of the tube housing: 9,000kJ (1,250kHU).
Cooling method passive or forced air and/or oil cooling.
Filtration: min 2.7 mm Al equivalent.
Floor Mounted Tube Stand:
Horizontal travel range: 138 cm to 220 cm, movement arrested by electromagnetic brakes;
Vertical travel range: 150 cm to 160 cm movement arrested by electromagnetic brakes;
Central beam height: min 35 cm to min 185 cm;
Rotation of tube around vertical axis: ± 90°; stop positions: 0°, ± 90°;
X-ray tube rotation: ± 120°; stop positions: 0°, ± 90°.
Flat Panel Detector:
Dimensions (active field): min 34 cm x max 43 cm;
Detector thickness: 16 mm
Pixel size: 139 µm;
Spatial resolution (Nyquist frequency): 3.6 lp/mm.
Collimator:
Inherent filtration: min 1 mm Al at 70 kV;
Full-field light localizer: 150 W halogen light, timer;
Built-in light switch should be provided.
Copper prefilter: Without filter, 0.1 mm, 0.2 mm, 0.3 mm Cu;
Rotation: Up to a maximum of ± 45°;
Collimation control: Manual.
Floor Mounted Bucky Stand:
Vertical movement range: min 150 cm;
Vertical travel range: min 35 cm to min 185 cm;
Grid Pb 10/80; f0 = 115 cm and 150 cm/180 cm, stationary grid;
Front-cover-to-detector distance: < 40 mm;
X-ray absorption: ≤ 1.2 mm Al eq at 100 kV;
Automatic exposure control: Standard.
Cassette size: accept all sizes from 5”x7” to 14”x17”.
Movement arrested by electromagnetic brakes.
Patient Table:
Table height: min 72 cm;
Tabletop width: min 80 cm;
Tabletop length: min 220 cm;
X-ray absorption: ≤ 1.2 mm Al eq at 100 kV;
Tabletop travel: mon two panes, ± 46 cm, longitudinal, ± 10 cm, transverse;
Max. patient weight: min 200 kg;
Grid Pb 10/80; f0 = 115 cm, stationary grid;
Max. patient coverage: min 185 cm without patient repositioning;
Detector cover range (edge-to-edge): ≥ 55 cm, movement arrested by electromagnetic brakes;
Tabletop – detector distance: ≤ 70 mm;
Workstation (Standard configuration):
The imaging station is specially designed for diagnostic radiology and delivers excellent image quality with a fast and seamless workflow;
High-resolution digital radiography system with DICOM network connection for image processing and image display on a preview monitor, min 2 MP;
High performance PC included;
DR printer
Accessories, spares and consumables
All standard accessories, consumables and parts required to operate the equipment.
Power supply:
Power supply: 415 ± 5%V (3 Phase), 50Hz
Environmental factors
The unit shall be capable of being stored continuously in ambient temperature of 0 -55deg C and relative humidity of 15-95%
12 | The unit shall be capable of operating in ambient temperature of 20-40 deg C and relative humidity of 80%
Standards & Safety Requirements
Should be CE or FDA approved product certificate.
Must submit ISO13485
Training
Must provide user & service training.
Warranty
Comprehensive warranty for 2 years.
Documentation
User (Operating) manual in English Should provide 2 sets (hardcopy and soft-copy)
Service (Technical / Maintenance) manual in English Should provide 2 sets (hardcopy and soft-copy)
List of important spare parts, accessories and consumables with their part numbers and costing.
15 | GENERAL RADIOGRAPHY (X-RAY) SYSTEM
1) Description of Function
General Radiography X-Ray System designed for routine diagnostic imaging of adult and pediatric patients with cost-effective configuration suitable for hospitals and medical centers with moderate workload.
The system shall include X-ray generator, X-ray tube, control console, radiographic table, wall bucky stand, and all essential accessories required for safe operation.
Compatible with conventional radiographic cassettes and Computed Radiography (CR) cassettes.
Parameter
Generator Power: ≥ 32 kW
Tube Voltage (kV): 40 – 125 kV, continuously adjustable
Tube Current (mA): ≥ 300 mA
Exposure Time: 5 ms – 5 seconds
Automatic Exposure Control (AEC): Optional or standard depending on manufacturer
X-Ray Tube: Dual focal spot
Small Focal Spot: ≤ 0.6 mm
Large Focal Spot: ≤ 1.2 mm
Tube Heat Capacity: ≥ 200 kHU
Radiographic Table: Floating tabletop with integrated bucky
Table Load Capacity: ≥ 180 kg
Wall Bucky Stand: Vertical movement, cassette compatible
Tube Stand: Floor-mounted type
Source to Image Distance (SID): Adjustable 100 – 180 cm minimum
Exposure Modes: Manual exposure and Anatomical Program Selection (APR)
Collimator: Manual collimation with light field indication
Dose Optimization: Basic exposure optimization system
2) Accessories & Consumables
The system shall be supplied with essential accessories only:
Wall Bucky Stand
Table Bucky
Radiographic Cassettes:
24 × 30 cm (2 pcs)
35 × 43 cm (2 pcs)
Anti-scatter Grid (10:1 ratio or higher)
Lead Aprons (2 pcs)
Thyroid Shield (1 pc)
Exposure Markers
Basic positioning aids
Cleaning kit
Spare fuses
3) Operating Environment
Ambient Temperature: 10°C – 40°C
Relative Humidity: 20% – 80% non-condensing
Power Supply: 220 – 240 VAC, 50 Hz
The system shall be suitable for continuous clinical use under standard hospital conditions.
4) Safety & Standards
The equipment shall comply with applicable international medical safety standards.
Parameter
Electrical Safety: IEC 60601-1
Radiography Equipment Safety: IEC 60601-2-54
Electromagnetic Compatibility: IEC 60601-1-2
Radiation Protection: Compliant with international radiation safety standards
Regulatory Approval: CE Marking or equivalent approval
5) Acceptance Criteria / Test Methodology
Acceptance testing shall be performed after installation and commissioning.
Parameter
Image Quality Test:
Verification using standard radiographic phantom.
kVp Accuracy Test:
Within manufacturer tolerance limits.
Exposure Time Accuracy Test:
Within manufacturer tolerance limits.
Mechanical Performance Test:
Smooth movement of table, tube stand, and bucky system.
Table Load Test:
Verification of safe operation at specified load.
Collimator Alignment Test:
Verification of light field and radiation field congruence.
Radiation Safety Test:
Compliance with applicable radiation safety limits.
6) User Training
The supplier shall provide on-site training for all operators and technical staff.
Training shall include:
Basic system operation
Patient positioning techniques
Exposure parameter selection
Radiation protection principles
Routine maintenance and cleaning
Basic troubleshooting
Training shall continue until staff demonstrate satisfactory operational competency.
7) Warranty
The system shall be covered by a minimum One (1) Year comprehensive warranty from the date of final acceptance.
Warranty shall include:
X-ray tube
Generator
Control console
Mechanical and electrical components
Labor and spare parts
Any defective parts shall be repaired or replaced at no additional cost.
8) After-Sales Service
The supplier shall provide essential after-sales support including:
Installation and commissioning
Preventive maintenance during warranty period
Corrective maintenance services
Technical support and assistance
Supply of genuine spare parts
Service shall be provided in accordance with manufacturer recommendations.
9) Installation & Commissioning
Installation and commissioning shall be performed by qualified personnel.
The supplier shall ensure:
Proper system installation
Functional verification of all components
Calibration and safety checks
Final operational readiness
Final acceptance shall be subject to successful completion of all tests.
10) Documentation
The supplier shall provide:
User Manual
Service Manual
Installation Manual
Spare Parts List
Basic Technical Data Sheets
Radiation Safety Guidelines
Documents shall be provided in printed and electronic formats.
11) Quality Certificates & Regulatory Compliance
The offered equipment shall be new, unused, and of current production.
The manufacturer shall provide valid documentation including:
CE Marking or equivalent
ISO 13485 Certificate
ISO 9001 Certificate
IEC 60601-1 Compliance
IEC 60601-2-54 Compliance
Declaration of Conformity
Certificate of Origin
All certificates shall be valid at the time of submission and delivery.