## Specifications
Technical specification - RFQ 8-2026
No | Technical Specification | Supplier offered specification | Compliant (yes/No) | Remarks / Deviation Details | Supporting Document
1 | Kidney Transplant Surgical Instrument Set
1. General Description
This specification covers a specialized surgical instrument set designed specifically for Kidney transplant procedure, featuring precision instruments required for vascular anastomosis urological reconstruction, and major abdominal surgery.
The instrument set is intended for use in operating rooms and surgical departments to support functions such as tissue dissection, vessel control, atraumatic clamping, surgical exposure, retraction, cutting, and suturing.
The set shall be suitable for repeated use under standard hospital sterilization conditions and shall ensure high precision, safety, and ergonomic handling during surgical procedures.
2. Material of Manufacture
All instruments shall be manufactured from high-quality surgical stainless steel (or equivalent medical-grade material) with the following characteristics:
High corrosion resistance.
Autoclavable and suitable for repeated sterilization cycles.
Non-toxic and biocompatible material.
High mechanical strength and durability
Smooth atraumatic finish to minimize tissue damage
Tungsten Carbide (TC) inserts for needle holders and cutting instruments where applicable
3. Technical Specifications (Instrument Categories & Sizes)
The set shall include instruments in the following categories. All dimensions and configurations are approximate and non-binding, and equivalent standard commercial sizes shall be accepted provided clinical function is maintained.
3.1 Atraumatic Vascular Clamps
Santulli atrauma Bulldog clamp : 70 mm (or equivalent)
Castaneda-Neonatal vascular clamps , delicate, (angled 30° and 45°): atrauma 12 cm (or equivalent)
Lambert-Kay atrauma clamp, curved, : approx. 21.5 cm and 25.5 cm (or equivalent)
Cooley atrauma clamp, S-shape, approx. 27.5 cm (or equivalent)
De Bakey-Mixter atrauma clamp, angled: approx. 28.5 cm (or equivalent)
3.2 Tissue Forceps
Duval tissue forceps, large jaw: approx. 20 cm (or equivalent)
Allis-Adair tissue forceps, straight, (10×11): approx. 16 cm (or equivalent)
DeBakey forceps: multiple lengths (16 cm, 20 cm, 24 cm, 30 cm or equivalent range)(or equivalent)
3.3 Retractors
Kocher retractor: approx. 40×10 mm, 22 cm (or equivalent)
Czerny double-ended retractor: approx. 17.5 cm (or equivalent)
Cushing hook: approx. 28 cm (or equivalent)
Desmarres vein retractor: blade approx. 16 mm, length 13 cm (or equivalent)
Kelly retractor: approx. 25×38 mm (or equivalent)
3.4 Scissors
Rib scissors (Giertz-Stille): approx. 27 cm (or equivalent)
Micro scissors (straight), Vannas cutting edge, Ø 8 mm,: approx. 18 cm (or equivalent)
Dean scissors, angled up, 17 cm (or equivalent)
Scissors ( curved, heavy-duty): 28 cm range (or equivalent)
Resano capsula scissors, angled, 24.5 cm blunt/blunt
Reynolds scissors, curved, 18 cm(or equivalent)
DeBakey arteriotomy scissors: approx. 17.5 cm (or equivalent)
3.5 Needle Holders
Micro needle holders with catch: 11*0.4 mm, Ø8mm, straight, 15cm (or equivalent)
Micro needle holders with catch: 11*0.4 mm, Ø9mm, straight, 18cm (or equivalent)
Micro needle holders with catch: 11*0.8 mm, Ø9mm, straight, 21 cm (or equivalent)
Ryder vascular TC needle holders: 13 cm, 15 cm, 18 cm, 23 cm (or equivalent)
3.6 Vascular Dilators
DeBakey dilators: Ø1.0 mm to Ø5.0 mm, approx. 19 cm (or equivalent graduated set)
3.7 Accessories
Finocchietto Rib spreader (adult), rigid lateral blades 73 x 65 mm: spreading 185 mm (or equivalent)
Kirk mallet: approx. 40 mm head, 750 g, 19 cm (or equivalent)
Kidney dishes, in sizes (170*100*35), (250*140*40) (or equivalent)
Round bowls, in sizes (Ø 40 mm, H 25 mm, 0.015 l), (Ø 255 mm, H 100 mm, 3.50 l), (Ø 260 mm, H 120 mm, 4.50 l) (or equivalent)
Sterilization cases in size (420*175*75) (or equivalent)
4. Quality Standards & Certifications
The supplier shall provide products compliant with the following international standards:
ISO 13485 (Medical Devices Quality Management System)
ISO 9001 (Quality Management System)
CE Mark (European Conformity) or equivalent international certification
FDA approval (if applicable based on origin)
RoHS compliance (where applicable)
Declaration of Conformity (DoC)
5. Warranty
Minimum warranty period: 24 months from date of acceptance
Warranty shall cover manufacturing defects and material failures
Includes replacement or repair of defective instruments
All replacement parts and labor shall be covered during warranty period
Response time for service support: ≤ 72 hours (recommended)
6. Preventive Maintenance & Service
The supplier shall provide:
Guidance for cleaning, sterilization, and maintenance procedures
Availability of spare parts or equivalent instruments for at least 10 years
7. Documentation (Mandatory)
The supplier must submit the following documents:
User, installation, and maintenance manuals
Technical datasheets for all instrument categories
Material composition certificates
Certificate of Origin
CE / ISO certificates
Sterilization and reprocessing guidelines
Declaration of conformity
2 | Hemodialysis Water Treatment Plant (HD-WTP)
Tender Document for 10 Dialysis Machines
1. General Description
The Contractor/Supplier shall design, manufacture, supply, install, test, commission, validate, and hand over a complete Hemodialysis Water Treatment Plant (HD-WTP) intended to produce ultrapure water suitable for hemodialysis applications.
The system shall be designed to serve a dialysis unit consisting of ten (10) hemodialysis machines with continuous, reliable, and safe operation in accordance with international dialysis water quality standards.
The supplied system shall include all equipment, accessories, piping, instrumentation, controls, electrical components, cooling systems, disinfection systems, and all necessary items required for complete and fully functional operation.
The HD-WTP shall be suitable for:
Conventional Hemodialysis (HD)
High-Flux Dialysis
Hemodiafiltration (HDF), where applicable
The complete system shall be designed for:
Continuous 24/7 recirculation
High microbiological safety
Ease of maintenance
Energy-efficient operation
Long-term operational reliability
2. Applicable Standards
The system shall comply with the latest editions of the following standards and guidelines:
Association for the Advancement of Medical Instrumentation Standards:
ANSI/AAMI RD52
ANSI/AAMI RD62
ANSI/AAMI/ISO 23500 Series
International Organization for Standardization:
ISO 23500 Parts 1–5
European Pharmacopoeia Requirements
World Health Organization Drinking Water Guidelines
CE MDR or equivalent international medical device certification
ISO 13485 certification for medical device manufacturing
3. System Capacity & Design Basis
The system shall be designed based on the following minimum criteria:
Number of dialysis machines :10 units
Average water consumption per machine :500–800 L/day
Estimated total daily demand :5–8 m³/day
Design safety margin :Minimum 30%
Total design capacity: 7–10 m³/day
RO production capacity : Minimum 500–700 L/hour
Operating duration : Minimum 16–20 hours/day
System operation : Continuous recirculation (24/7)
The Supplier shall verify all design calculations based on actual site conditions and raw water analysis.
4. Feed Water Requirements
The Supplier shall review available raw water analysis prior to final design.
The following parameters shall be considered during system design:
Hardness
Iron
Manganese
TDS
Turbidity
Free chlorine/chloramines
Silica
SDI
Microbiological contamination
If no water analysis is available, the Supplier shall clearly state all assumed feed water conditions.
5. Final Water Quality Requirements
The HD-WTP shall continuously achieve dialysis water quality compliant with AAMI/ISO standards.
Product Water Quality
Bacteria :< 0.1 CFU/mL
Endotoxins: < 0.03 EU/mL
Free chlorine: 0 mg/L
TOC :< 0.5 mg/L
Product water conductivity :≤ 5 µS/cm preferred
Alarm conductivity: >10 µS/cm
Online Monitoring (Mandatory)
Continuous online monitoring shall include:
Conductivity
Pressure
Flow rate
Temperature
Free chlorine
Tank level
Loop return conductivity
Monitoring locations shall include:
Post-RO Pass 1
Post-RO Pass 2
Distribution loop return
6. Pretreatment System
The pretreatment section shall include all necessary stages required to protect the RO membranes and ensure reliable operation.
6.1 Primary Storage Tank
The storage tank shall be sanitary SS316L construction.
Requirements
Capacity: 3000–5000 liters
0.2 µm hydrophobic vent filter
Overflow protection
SS316L high-pressure pumps
6.2 Multimedia Sand Filter
The multimedia filter shall provide high-efficiency suspended solids and turbidity removal.
Requirements
Automatic backwash operation
Differential pressure monitoring
FRP or SS316L pressure vessel
Automatic control valve
Sampling ports
6.3 Iron Removal Filtration System (Mandatory)
A dedicated iron removal system shall be provided as a mandatory pretreatment stage.
Performance Requirements
Outlet iron concentration:<0.05 mg/L preferred
Maximum allowable :≤0.1 mg/L
Media
Acceptable media include:
Birm
Manganese Greensand
Katalox Light
Equivalent catalytic media
Configuration
FRP or SS316L pressure vessel
Fully automatic operation
Automatic backwash
PLC integration
Sampling ports (inlet/outlet)
Installation Location
The iron removal filter shall be installed:
After multimedia filtration
Before activated carbon filtration
6.4 Activated Carbon Filtration System (Safety Critical Stage)
A catalytic activated carbon filtration system shall be provided for complete removal of chlorine and chloramines.
Configuration
Single Vessel (subject to approval)
Requirements
Catalytic activated carbon media
EBCT ≥10 minutes at peak flow
Automatic backwash
Proper internal distribution
Mandatory Safety Features
Online free chlorine analyzer
Continuous monitoring
Audible and visual alarms
PLC integration
Performance Requirement
Outlet free chlorine concentration:
0 mg/L at all times
6.5 Water Softener
The softener system shall include:
Duplex configuration
Automatic regeneration
Continuous operation
Online hardness monitoring
Performance
Residual hardness:
<1 ppm as CaCO₃
RO shutdown shall occur upon hardness breakthrough.
6.6 Chemical Dosing System (Optional)
Where required, anti-scalant dosing shall be provided.
The dosing system shall:
Be interlocked with RO operation
Include chemical dosing pump
Include dosing tank and level monitoring
7. Reverse Osmosis (RO) System
Configuration
Single Pass RO System
Capacity
Minimum production:
500–700 L/hour
General Requirements
Membrane rejection ≥99% (with size (40*40))
SS316L high-pressure pumps
Automatic flushing system
CIP system
Membrane integrity testing
Performance
Recovery rate :60–75%
SDI at RO inlet:<3
8. Distribution System
Distribution Loop
General Requirements
Continuous recirculation (24/7)
Ring loop configuration
SS316L electropolished piping
Hygienic sanitary fittings
No dead legs permitted
Design Criteria
Velocity:1.0–1.5 m/s
Loop turnover:6–10 times/hour
Compatibility
The loop shall be compatible with:
Thermal disinfection (80–85°C)
Chemical disinfection
9. Chiller System (Mandatory)
A dedicated industrial chiller system shall be provided for cooling the dialysis water distribution loop and storage system to minimize microbiological growth.
Performance Requirements
Loop temperature :20–25°C adjustable
Maximum allowable temperature :≤25°C
Ambient design condition :45–50°C
Chiller Components
The system shall include:
Air-cooled chiller
SS316L plate heat exchanger
Circulation pump
Expansion tank
Digital temperature controller
Monitoring & Control
The system shall include:
Online temperature monitoring
High-temperature alarm
PLC integration
Automatic operation
Thermal Disinfection Compatibility
The chiller shall automatically isolate or bypass during thermal disinfection cycles.
10. Disinfection System
The system shall include automated disinfection capability for:
RO system
Distribution loop
Methods
Preferred:
Thermal disinfection
Chemical disinfection:
Peracetic acid
Citric acid
Features
Automated cycles
Validation and logging
Safety interlocks
11. Control & Automation System
The system shall include a PLC-based control system with HMI touchscreen interface.
Monitoring Parameters
Pressure
Flow
Conductivity
Temperature
Free chlorine
Chiller status
Alarm status
Features
Alarm and fault logging
Data logging (minimum 12 months)
Temperature trending
Optional remote monitoring
User access control
12. Electrical System
Requirements
Power supply:
220–240V / 380–415V
50–60 Hz
MCC panel
Surge protection
UPS backup:
Minimum 60 minutes
13. Materials of Construction
Wetted Parts
Mandatory materials:
SS316L stainless steel
PVDF
Medical-grade PVC
Gaskets
EPDM
PTFE
Not Permitted
Galvanized steel
Industrial PVC
Low-grade plastics
Carbon steel in wetted areas
14. Compliance & Validation
The Supplier shall provide:
IQ / OQ / PQ validation
FAT
SAT
Water quality validation reports
Temperature validation
Continuous operation trial
Compliance Standards
ISO 13485
CE MDR (or equivalent)
AAMI / ISO 23500
Pharmacopoeia standards
15. Documentation Requirements
The Supplier shall provide:
PFD
P&ID
General Arrangement Drawings
Electrical Schematics
O&M Manuals
Spare Parts List
Validation Reports
PLC/SCADA Backup
Chiller calculations
As-built drawings
16. Warranty
Minimum Warranty Period
24 months from commissioning
Coverage
Mechanical components
Electrical components
Pumps
Instrumentation
PLC/control system
Conditions
Repair or replacement
Response time: 48–72 hours
17. Training Requirements
The Supplier shall provide on-site training covering:
Operation
Monitoring
Maintenance
Disinfection
Troubleshooting
Training Duration
Minimum 3–5 days
18. After-Sales Service
The Supplier shall provide:
Technical support
Emergency support
Spare parts availability for minimum 10 years
Optional Annual Maintenance Contract (AMC)
19. Vendor Qualification Requirements
The bidder shall demonstrate:
Experience with similar dialysis water systems
Availability of local technical support
Availability of spare parts and consumables
Qualified service personnel