## Specs - 9-2026
Technical specification - RFQ 9-2026
No | Technical Specification | Supplier offered specification | Compliant (yes/No) | Remarks / Deviation Details | Supporting Document
1 | Inverted Fluorescence Microscope
1. General Description
A reliable inverted fluorescence microscope designed for routine laboratory applications, including:
Cell culture observation (including kidney cells)
Basic immunofluorescence applications
Evaluation of sample uniformity and consistency
Routine imaging and documentation
2. Optical System
Infinity-corrected optical system
Standard diffraction-limited optical resolution
Field of view ≥ 20–22 mm
Inverted configuration (objectives positioned below the stage)
3. Observation Modes
Brightfield
Phase Contrast
Fluorescence (limited to standard channels)
Differential Interference Contrast (DIC) (optional)
4. Objectives
Plan Achromat objectives (Plan Apochromat optional)
Magnifications: 4×, 10×, 20×, 40× (60× optional)
Numerical aperture suitable for routine imaging (up to approximately 1.3)
5. Fluorescence System
LED illumination source with long lifetime (≥ 30,000–50,000 hours)
Support for standard fluorescence channels (e.g., DAPI, FITC, TRITC)
Filter cubes with appropriate dichroic mirrors
Manual filter selection (motorized optional)
6. Stage and Focus System
Mechanical XY stage
Compatible with culture dishes and multiwell plates
Coarse and fine focusing mechanisms
Stable and precise manual focusing system suitable for routine use
7. Imaging System
Digital camera (CMOS or CCD)
Resolution ≥ 5 megapixels
Camera with high sensitivity and low noise suitable for fluorescence imaging
Global shutter or rolling shutter with fast readout capability suitable for live imaging applications
Minimum exposure time ≤ 1 ms
Live image display capability
Image capture and storage functionality
Basic image acquisition features
8. Environmental Conditions
Operation under standard laboratory conditions
No live-cell environmental control required
9. Software
Image acquisition software
Multi-channel fluorescence imaging support
Basic image analysis (measurement, annotation, optional cell counting)
Compatible with standard operating systems
10. Electrical Requirements
Power supply: 220–240 V AC, 50/60 Hz
11. Safety Requirements
UV protection shielding
Automatic light shut-off feature
Compliance with IEC 61010 or equivalent safety standards
12. Quality and Compliance
ISO 9001 certification
CE marking
RoHS compliance
ISO 13485 and FDA (if applicable)
13. Documentation (Mandatory)
The supplier shall provide:
User manual
Installation manual
Basic service documentation
Technical datasheet
Calibration information
Copies of relevant certificates
14. Warranty (Mandatory)
Minimum warranty period: 24 months from the date of installation
Warranty shall include:
All parts and labor
On-site service and support
Replacement of defective components
Response time ≤ 72 hours
15. Training (Mandatory)
On-site training (1–2 days) covering:
End users (operation and imaging)
Technical staff (basic maintenance)
Training materials must be provided
16. Preventive Maintenance and Service
Basic preventive maintenance schedule
Availability of spare parts for at least 5 years
Maintenance records recommended
Optional Annual Maintenance Contract (AMC)
17. Accessories
Computer workstation
Calibration slide
Standard fluorescence filter sets