## Sheet1
Technical Datasheet
No | Technical Specification | Supplier offered specification | Compliant (yes/No) | Remarks / Deviation Details | Supporting Document
1 | Digital X-ray Machine
Description of Function
A general purpose digital X-ray machine for routine X-ray examinations at healthcare facilities.
Operational Requirements
It shall be suitable to be used for adult and pediatric patients in general Radiography examination.
System Configuration
X-ray Generator,1 unit
X-Ray tube & tube support system, 1 unit
Radiographic patient table, fixed height1 unit
Floor mounted bucky stand, 1 unit
Technical Specifications
Should be floor to floor mounted
X-ray Generator:
Microprocessor based, high frequency inverter generator, the generator shall have at least 100kHz.
Generator Output: not less than 30 Kw
Radiographic voltage: 40 kV to 140kV, in 1Kv step.
Radiographic current approx., 10 to 500 mA
mAs range: 0.5 mAs to 650 mAs.
Anatomical Programmable Radiographic mode shall be available.
Manual & automatic exposure control and automatic brightness control shall be available.
Shall come with overload protection device.
X-Ray Tube:
Maximum tube output shall match with the generator output
X-ray tube rotating: +/-90°.
Anode rotating speed: 2700 rpm or more.
Large focus not more than 1.2 mm.
Small focus not more than 0.6 mm.
Maximum tube voltage 140 KV.
Max housing heat dissipation rate: 250HU/s
Anode heat storage capacity: min 140 kHU (100kJ);
Max. heat storage capacity of the tube housing: 9,000kJ (1,250kHU).
Cooling method passive or forced air and/or oil cooling.
Filtration: min 2.7 mm Al equivalent.
Floor Mounted Tube Stand:
Horizontal travel range: 138 cm to 220 cm, movement arrested by electromagnetic brakes;
Vertical travel range: 150 cm to 160 cm movement arrested by electromagnetic brakes;
Central beam height: min 35 cm to min 185 cm;
Rotation of tube around vertical axis: ± 90°; stop positions: 0°, ± 90°;
X-ray tube rotation: ± 120°; stop positions: 0°, ± 90°.
Flat Panel Detector:
Dimensions (active field): min 34 cm x max 43 cm;
Detector thickness: 16 mm
Pixel size: 139 µm;
Spatial resolution (Nyquist frequency): 3.6 lp/mm.
Collimator:
Inherent filtration: min 1 mm Al at 70 kV;
Full-field light localizer: 150 W halogen light, timer;
Built-in light switch should be provided.
Copper prefilter: Without filter, 0.1 mm, 0.2 mm, 0.3 mm Cu;
Rotation: Up to a maximum of ± 45°;
Collimation control: Manual.
Floor Mounted Bucky Stand:
Vertical movement range: min 150 cm;
Vertical travel range: min 35 cm to min 185 cm;
Grid Pb 10/80; f0 = 115 cm and 150 cm/180 cm, stationary grid;
Front-cover-to-detector distance: < 40 mm;
X-ray absorption: ≤ 1.2 mm Al eq at 100 kV;
Automatic exposure control: Standard.
Cassette size: accept all sizes from 5”x7” to 14”x17”.
Movement arrested by electromagnetic brakes.
Patient Table:
Table height: min 72 cm;
Tabletop width: min 80 cm;
Tabletop length: min 220 cm;
X-ray absorption: ≤ 1.2 mm Al eq at 100 kV;
Tabletop travel: mon two panes, ± 46 cm, longitudinal, ± 10 cm, transverse;
Max. patient weight: min 200 kg;
Grid Pb 10/80; f0 = 115 cm, stationary grid;
Max. patient coverage: min 185 cm without patient repositioning;
Detector cover range (edge-to-edge): ≥ 55 cm, movement arrested by electromagnetic brakes;
Tabletop – detector distance: ≤ 70 mm;
Workstation (Standard configuration):
The imaging station is specially designed for diagnostic radiology and delivers excellent image quality with a fast and seamless workflow;
High-resolution digital radiography system with DICOM network connection for image processing and image display on a preview monitor, min 2 MP;
High performance PC included;
DR printer
Accessories, spares and consumables
All standard accessories, consumables and parts required to operate the equipment.
Power supply:
Power supply: 415 ± 5%V (3 Phase), 50Hz
Environmental factors
The unit shall be capable of being stored continuously in ambient temperature of 0 -55deg C and relative humidity of 15-95%
The unit shall be capable of operating in ambient temperature of 20-40 deg C and relative humidity of 80%
Standards & Safety Requirements
Should be CE or FDA approved product certificate.
Must submit ISO13485
Training
Must provide user & service training.
Warranty
Comprehensive warranty for 2 years.
Documentation
User (Operating) manual in English Should provide 2 sets (hardcopy and soft-copy)
Service (Technical / Maintenance) manual in English Should provide 2 sets (hardcopy and soft-copy)
List of important spare parts, accessories and consumables with their part numbers and costing.