## Specs -12-2026
Technical Datasheet - RFQ - 12-2026
No | Technical Specification | Supplier offered specification | Compliant (yes/No) | Remarks / Deviation Details | Supporting Document
1 | ENT Diagnostic Endoscopy System
Basic Diagnostic Endoscopy System for ENT
1. General Description
The ENT Diagnostic Endoscopy System shall be a complete diagnostic endoscopy system designed for otorhinolaryngology (ENT) examinations of adult and pediatric patients.
The system shall be compatible for routine diagnostic examination of:
Nasal cavity.
Nasopharynx.
Pharynx.
Larynx.
Other accessible ENT anatomical areas as applicable.
The system shall provide clear, high-quality visualization and adequate illumination for routine diagnostic ENT examinations.
The system is intended for diagnostic use only and shall not be required to support operative or surgical endoscopic procedures.
2. Scope of Supply
The system shall include, as a minimum:
2.1 Camera and Image Processing System
Full HD camera system with minimum resolution of 1920 × 1080.
Camera head compatible with the supplied flexible diagnostic endoscope.
Image processor/control unit.
Basic image adjustment functions.
Digital zoom function.
Suitable video output for connection to the supplied medical monitor.
Clear and stable image suitable for diagnostic ENT examinations.
2.2 Illumination System
Integrated LED cold-light source, maximum up to 150 watt .
Adjustable illumination intensity.
Compatible light guide/fiber-optic cable for the supplied flexible endoscope, diameter 3.5 mm .
Adequate illumination for routine diagnostic ENT examinations.
Long-life LED light source.
3. Medical Monitor
The system shall include:
One medical-grade monitor.
Minimum screen size: 24 inches.
Full HD resolution (1920 × 1080) or higher.
Suitable for displaying live endoscopic images.
Compatible with the supplied camera/image processor.
Additional monitors are not required in the base configuration.
4. Diagnostic ENT Telescope
The system shall include Rigid Telescope for Adult diagnostic ENT and Telescope pediatric use.
Telescope 30°, enlarged view, diameter 4 mm, length 18 cm
Telescope 45°, enlarged view, diameter 4 mm, length 18 cm
Telescope 30°, enlarged view, diameter 2.7 mm, length 18 cm
Telescope 0°, enlarged view, diameter 2.7 mm, length 18 cm
The endoscope shall be suitable for routine diagnostic examination of:
Nasal cavity.
Nasopharynx.
Pharynx.
Larynx.
Minimum requirements:
Compatible with the supplied camera and image processor.
Suitable for adult diagnostic examinations.
Suitable for pediatric examinations where applicable.
High-quality image suitable for diagnostic assessment.
5. Mobile Equipment Tower
The system shall include a basic mobile tower suitable for housing the supplied equipment.
Minimum requirements:
Stable mobile frame.
Lockable castors.
Shelves for the camera/image processor, light source and monitor.
Basic cable management.
Adequate ventilation.
Stable construction with anti-tip protection.
6. Infection Control and Reprocessing
The endoscope shall be compatible with the manufacturer's recommended cleaning, disinfection and reprocessing procedures.
All patient-contact components shall be suitable for the applicable hospital infection-control procedures.
External equipment surfaces shall be smooth and easy to clean and disinfect.
The supplier shall provide the manufacturer's instructions for cleaning, disinfection and reprocessing.
7. Safety and Reliability
The system shall include:
Electrical safety in accordance with applicable medical electrical equipment standards.
Overheating protection for the LED light source.
Stable mobile tower with lockable castors.
Anti-tip or equivalent stability protection.
Appropriate optical and electrical safety features.
Safe operation of the camera, light source and endoscope.
8. Standards and Regulatory Compliance
The supplied equipment shall comply, where applicable, with:
ISO 13485 or equivalent medical-device quality management requirements.
IEC 60601-1 – General requirements for basic safety and essential performance of medical electrical equipment.
Applicable requirements for endoscopic equipment.
Applicable CE Marking and/or other relevant regulatory approval.
The supplier shall provide appropriate certificates and declarations of conformity for the supplied equipment.
9. Documentation
The supplier shall provide:
User Manual.
Service Manual or applicable technical/service documentation.
Technical Datasheet/Catalog.
Certificate of Conformity.
CE documentation or applicable regulatory approval.
Packing List.
Warranty Certificate.
Manufacturer's cleaning, disinfection and reprocessing instructions for the flexible endoscope.
10. Warranty and After-Sales Service
Minimum warranty: 24 months from the date of acceptance/commissioning.
Availability of spare parts and technical support for 10 years.
Technical assistance for maintenance and troubleshooting.
Preventive maintenance recommendations.
Availability of replacement/reprocessing accessories for the flexible endoscope.
11. Training
The supplier shall provide on-site training for relevant hospital personnel, including:
ENT physicians.
Nurses/clinical staff.
Biomedical/technical staff responsible for the equipment.
Training shall cover, as applicable:
Basic operation of the ENT diagnostic endoscopy system.
Operation of the camera and image processor.
Operation and adjustment of the LED light source.
Basic troubleshooting.
Cleaning and disinfection of the endoscope.
Reprocessing and sterilization procedures in accordance with the manufacturer's instructions.
Routine care and preventive maintenance.
Basic safety precautions.
Proper storage and handling of the endoscope.
The supplier shall provide at least one on-site training session following installation and commissioning.
Training shall be conducted by qualified personnel from the supplier or manufacturer.
Training shall be included in the base supply price and shall not be charged separately.
2 | SLIT-LAMP BIOMICROSCOPE WITH APPLANATION TONOMETER
2.1 Technical Requirements
1. Binocular stereoscopic viewing system.
2. High-resolution optical system.
3. Adjustable interpupillary distance.
4. Minimum five magnification levels or continuous magnification.
5. Magnification approximately 6×–40× or higher.
6. Wide field of view.
7. Adjustable eyepieces.
8. Joystick-controlled movement.
9. Smooth X/Y/Z movement.
10. Adjustable chin and forehead rests.
11. Stable patient positioning.
12. LED illumination preferred.
13. Adjustable illumination intensity.
14. Adjustable slit width, height, angle and orientation.
15. Diffuse illumination capability.
2.2 Applanation Tonometer
1. Goldmann-type applanation tonometer.
2. Integrated or fully compatible with the slit lamp.
3. Suitable for routine intraocular pressure measurement.
4. Compliant with ISO 8612:2009 or the current applicable edition.
2.3 Standards
1. ISO 10939:2017 or current applicable edition.
2. Applicable IEC 60601 requirements.
2.4 Electrical and Environmental Requirements
1. Power supply: 220 V ±10%, 50 Hz, unless the manufacturer specifies a compatible equivalent input.
2. Supplier shall state power consumption.
3. Supplier shall state operating temperature range.
4. Supplier shall state operating relative humidity range.
5. Equipment shall be suitable for the environmental conditions of Khartoum, Sudan.
6. Suitable electrical protection and proper protective earthing shall be provided where applicable.
2.5 Required Documents and Quality Certificates
1. Valid ISO 13485:2016 certificate for the manufacturer.
2. Applicable regulatory approval/certification.
3. Declaration of Conformity.
4. Applicable IEC/ISO certificates or test reports.
5. Manufacturer's official technical datasheet/catalogue.
6. User manual.
7. Service/maintenance manual where applicable.
8. Manufacturer Authorization Letter.
9. Calibration certificate or calibration procedure where applicable.
10. Country of origin and manufacturing location.
11. Exact manufacturer and model number.
2.6 Installation and Commissioning
1. Delivery.
2. Installation and assembly.
3. Configuration.
4. Calibration where applicable.
5. Functional testing.
6. Demonstration.
7. Final acceptance testing.
2.7 Training
1. On-site training for ophthalmologists.
2. Training for nurses and technicians.
3. Basic maintenance and troubleshooting training for biomedical engineers.
4. Cleaning and disinfection/reprocessing training where applicable.
2.8 Warranty, Spare Parts and Service
1. Minimum 24-month warranty from successful installation and final acceptance.
2. Warranty shall cover the main equipment and standard manufacturer-supplied accessories.
3. Technical support during the warranty period.
4. Preventive maintenance where applicable.
5. Genuine spare parts shall be available for at least 10 years from final acceptance.
6. Supplier shall state local/regional service arrangements.
7. Supplier shall state expected service response time.
8. Supplier shall state technical support location.
3 | AUTO-REFRACTOMETER / AUTO-KERATOMETER
3.1 Measurements
1. Sphere.
2. Cylinder.
3. Axis.
4. Spherical equivalent.
5. K1/K2.
6. Corneal curvature.
7. Corneal astigmatism and axis.
3.2 Features
1. Automatic or assisted alignment.
2. Automatic measurement.
3. Multiple readings and averaging.
4. Right/left eye identification.
5. Adjustable chin and forehead rests.
6. Digital display.
7. Patient data storage.
8. Printer or digital data output preferred.
3.3 Electrical and Environmental Requirements
1. Power supply: 220 V ±10%, 50 Hz, where applicable.
2. Supplier shall state power consumption.
3. Supplier shall state operating temperature.
4. Supplier shall state operating relative humidity.
3.4 Quality and Documents
1. ISO 13485:2016.
2. Applicable regulatory approval.
3. Declaration of Conformity.
4. Applicable IEC/ISO documentation.
5. Manufacturer datasheet/catalogue.
6. User manual.
7. Service manual where applicable.
8. Calibration documentation.
9. Manufacturer Authorization Letter.
3.5 Installation, Training and Warranty
1. Installation and commissioning.
2. Calibration.
3. Functional testing.
4. On-site training.
5. Minimum 24-month warranty.
6. Minimum 10-year spare-parts availability.
7. Technical support.
4 | Nd:YAG OPHTHALMIC LASER
4.1 Applications
1. Posterior capsulotomy.
2. Peripheral iridotomy.
3. Other clinically indicated photodisruptive procedures.
4.2 Laser
1. Wavelength: 1064 nm.
2. Adjustable energy.
3. Single/multiple pulse capability.
4. Stable aiming beam.
5. Accurate focusing.
4.3 Delivery
1. Integrated or slit-lamp mounted.
2. Binocular viewing.
3. Joystick control.
4. Footswitch.
4.4 Safety
1. Key switch.
2. Emergency shutdown.
3. Safety interlock.
4. Emission indicator.
5. Laser warning labels.
4.5 Standard
1. IEC 60601-2-22:2019 + AMD1:2026, where applicable.
4.6 Electrical and Environmental Requirements
1. 220 V ±10%, 50 Hz.
2. Supplier shall state power consumption.
3. Operating temperature and humidity shall be stated.
4. Electrical protection shall be provided.
4.7 Documents
1. ISO 13485:2016.
2. Applicable regulatory approval.
3. Declaration of Conformity.
4. Applicable IEC laser safety documentation.
5. Calibration documentation.
6. Datasheet/catalogue.
7. User manual.
8. Service manual.
9. Manufacturer Authorization Letter.
4.8 Installation, Training and Warranty
1. Installation.
2. Calibration.
3. Acceptance testing.
4. Clinical training.
5. Laser safety training.
6. Biomedical engineering training.
7. Minimum 24-month warranty.
8. Minimum 10-year spare-parts availability.
9. After-sales technical support.
5 | Digital Electroencephalography (EEG) System
Technical Specification
1. General Description
The Supplier shall supply, install, test, commission, and hand over a complete Digital EEG System suitable for routine adult and pediatric EEG examinations.
The system shall include all equipment and accessories required for complete routine EEG operation.
2. Applicable Standards
The system shall comply with the latest applicable editions of:
IEC 80601-2-26 – Electroencephalographs
IEC 60601-1 – Medical electrical safety
IEC 60601-1-2 – Electromagnetic compatibility
International 10–20 EEG Electrode System
CE marking or equivalent recognized regulatory approval
3. EEG Acquisition System
The EEG acquisition system shall meet the following minimum requirements:
EEG channels: Minimum 21 simultaneous channels
ECG: Minimum 1 channel
EOG: Minimum 2 channels
Sampling rate: Minimum 256 Hz
A/D resolution: Minimum 16 bit
Input impedance: Minimum 100 MΩ
CMRR: Minimum 100 dB
EEG bandwidth: 0.1–70 Hz or wider
Recording: Continuous digital recording
Electrode impedance: Individual measurement and display
Patient isolation: Required
4. Electrode System
The system shall include:
Complete adult EEG electrode set
Pediatric EEG electrode accessories
Reference electrodes
Ground electrode
Electrode cables
Electrode paste/gel
The electrodes shall be compatible with the International 10–20 System.
5. EEG Software
The software shall provide:
EEG recording
Real-time waveform display
Playback and review
Bipolar montage
Referential montage
User-defined montage
Adjustable sensitivity
Adjustable time scale
High-pass filter
Low-pass filter
50 Hz notch filter
Electrode impedance display
Event marking
Patient annotation
Patient data storage
EEG report generation
PDF report export
6. Photic Stimulation
The system shall include a photic stimulator with:
Frequency range: 1–30 Hz
Manual operation
Automatic frequency selection
Synchronized event marking
7. Hyperventilation
The software shall provide:
Hyperventilation event marking
Timer
Synchronized recording
8. ECG & EOG
The system shall include:
1 ECG channel
2 EOG channels
Required patient cables
9. Data Storage & Export
The system shall provide:
Patient database
Patient identification
Recording storage
Data backup
EDF/EDF+ export
PDF report export
10. Computer Workstation
The supplier shall provide a computer workstation suitable for the supplied EEG software, with minimum:
Processor: Intel Core i5 or equivalent
RAM: 8 GB
SSD: 512 GB
Monitor: 21 inches or larger
Keyboard and mouse
Compatible operating system
Licensed EEG software
11. UPS
The supplier shall provide a UPS for the complete EEG system.
Minimum requirements:
220–240 V / 50 Hz
Voltage regulation
Surge protection
Backup time: ≥30 minutes
12. Accessories
The base supply shall include all accessories required for routine EEG operation, including:
EEG amplifier/headbox
EEG electrodes
Reference and ground electrodes
ECG/EOG cables
Photic stimulator
Computer
Monitor
UPS
All necessary connection and power cables
13. Installation & Commissioning
The Supplier shall provide:
Installation
System configuration
Functional testing
Testing of all EEG channels
Impedance testing
Photic stimulation testing
Electrical safety testing
Commissioning
14. Training
The Supplier shall provide Abroad training covering:
System operation
10–20 electrode placement
Patient preparation
EEG recording
Impedance measurement
Montages
Filters
Photic stimulation
Basic troubleshooting
Minimum training: 2 working days
15. Warranty
Minimum warranty: 24 months from commissioning
The warranty shall cover the main EEG system, amplifier/headbox, computer, monitor, photic stimulator, UPS, and supplied accessories.
Technical response time: ≤72 hours
16. Documentation
The Supplier shall provide:
User manual
Technical specifications
Regulatory certificate
Calibration/test certificate
Warranty certificate
Basic maintenance instructions
Spare parts list
17. Vendor Qualification
The bidder shall demonstrate:
Manufacturer authorization
Previous experience supplying EEG systems
Availability of technical support
Availability of spare parts
6 | Electroconvulsive Therapy (ECT) System
Tender Technical Specification – Tigani El Mahi Hospital
1. General Description
The Supplier shall supply, install, test, commission, and hand over a complete Electroconvulsive Therapy (ECT) System suitable for routine hospital ECT procedures.
The system shall include the ECT treatment unit, patient cables, EEG monitoring, ECG monitoring, stimulus control, safety features, accessories, and all items required for complete operation.
2. Applicable Standards
The system shall comply with applicable:
IEC 60601-1 – Medical electrical safety
IEC 60601-1-2 – Electromagnetic compatibility
Applicable standards for ECT equipment
ISO 13485 – Medical device quality management
CE marking or equivalent recognized regulatory approval
3. ECT Treatment Unit
The ECT unit shall provide:
Controlled electrical stimulus
Adjustable stimulus parameters
Constant-current output
Bilateral and unilateral electrode placement
Digital display of treatment parameters
Treatment parameter storage
Automatic calculation/display of delivered charge
Emergency stop function
4. Stimulus Parameters
The system shall provide adjustable:
Pulse width
Frequency
Pulse number/duration
Current
Stimulus charge
Minimum requirements:
Maximum current: ≥800 mA
Pulse width: Adjustable
Frequency: Adjustable
Delivered charge: Displayed digitally
The bidder shall clearly state the complete stimulus parameter ranges.
5. EEG Monitoring
The system shall include integrated EEG monitoring for monitoring seizure activity.
Minimum requirements:
Minimum 2-channel EEG
Real-time EEG waveform display
Seizure duration measurement
Digital recording
Event marking
Adjustable EEG sensitivity
Basic EEG filters
6. ECG Monitoring
The system shall provide:
ECG monitoring
Minimum 1 ECG channel
Heart rate display
ECG waveform display
Patient cable
7. Patient Monitoring
The ECT system shall provide or be compatible with routine physiological monitoring, including:
ECG
Heart rate
SpO₂
Non-invasive blood pressure (NIBP)
If the hospital already has a compatible patient monitor, the bidder may quote the monitor separately.
8. Safety Features
The system shall include:
Emergency stop
Patient isolation
Electrical safety protection
Automatic discharge protection
Fault indication
Treatment lockout where applicable
Audible and visual alarms
Self-test at startup
9. Treatment Recording
The system shall provide:
Patient identification
Treatment date/time
Stimulus parameters
Delivered charge
EEG seizure recording
Seizure duration
Treatment record storage
Report printing/export
10. Electrodes & Accessories
The initial supply shall include:
Adult ECT electrode set
Patient treatment cables
EEG cables
ECG cable
Ground/reference accessories
Electrode contact material
Required connection cables
Reusable electrodes are preferred where practical.
11. Computer / Data Management
A dedicated computer shall not be mandatory if the ECT unit has an integrated recording and reporting system.
If a computer is required by the offered system, it shall be quoted separately.
12. Electrical Requirements
The system shall operate on:
220–240 V AC
50 Hz
The system shall include appropriate electrical protection.
13. UPS
A UPS shall be provided for the ECT system.
Minimum backup:
≥30 minutes
The UPS shall provide:
Voltage regulation
Surge protection
Overload protection
14. Installation & Commissioning
The Supplier shall provide:
Installation
Functional testing
Stimulus output verification
EEG testing
ECG testing
Safety testing
Alarm testing
Electrical safety testing
Final commissioning
15. Training
The Supplier shall provide on-site training for psychiatrists, anesthesia staff, nurses, and biomedical personnel covering:
System operation
Patient connection
Electrode placement
Stimulus parameter setting
EEG monitoring
ECG monitoring
Safety procedures
Cleaning
Basic troubleshooting
Minimum training:
2 working days
16. Warranty
Minimum warranty:
24 months from commissioning
Warranty shall cover:
ECT treatment unit
EEG module
ECG module
Patient cables
Electrodes
Footswitch/control accessories
Manufacturer-supplied accessories
Technical response time:
≤72 hours
17. Documentation
The Supplier shall provide:
User manual
Technical manual
Service/maintenance instructions
Electrical safety certificate
Regulatory certificate
Calibration certificate
Warranty certificate
Spare parts list
Recommended consumables list
18. Vendor Qualification
The bidder shall demonstrate:
Manufacturer authorization
Previous experience supplying ECT systems
Hospital installation references
Technical support availability
Spare parts availability
Calibration/service capability
7 | Electrical Nerve & Muscle Stimulator
Tender Technical Specification – Tigani El Mahi Hospital
1. General Description
The Supplier shall supply, install, test, and commission a complete Electrical Nerve and Muscle Stimulator suitable for routine physiotherapy and rehabilitation applications.
The system shall be suitable for:
TENS
EMS/NMES
Muscle stimulation
Pain management
Muscle re-education
2. Stimulator Unit
The stimulator shall provide:
Minimum 2 independent channels
Digital display
Adjustable intensity for each channel
Independent channel control
Continuous and pulsed stimulation
Timer
Patient safety control
Audible/visual indication of operation
3. Stimulation Parameters
The system shall provide adjustable:
Pulse frequency: approximately 1–150 Hz or wider
Pulse width: approximately 50–400 µs or wider
Intensity: independently adjustable for each channel
Treatment time: adjustable up to 60 minutes
Continuous and intermittent stimulation modes
The bidder shall state the exact parameter ranges.
4. Treatment Modes
The system shall provide, as a minimum:
TENS mode
EMS/NMES mode
Continuous stimulation
Burst/intermittent stimulation
Adjustable treatment timer
Pre-programmed clinical protocols shall not be mandatory.
5. Patient Connection & Electrodes
The system shall include:
Minimum 4 reusable electrodes
Patient cables
Electrode leads
Connection accessories
Electrodes shall be suitable for routine physiotherapy use and shall be replaceable.
6. Safety Features
The system shall include:
Patient current limitation
Output isolation
Automatic output reduction to zero when treatment is stopped
Open-circuit indication
Overload protection
Emergency stop/output interruption
Audible/visual fault indication
7. Power Supply
The system shall operate on:
220–240 V AC
50 Hz
The supplier shall provide the required power adapter and electrical accessories.
8. Accessories
The base supply shall include:
Stimulator unit
Patient cables
Minimum 4 electrodes
Power cable/adapter
Operating manual
Basic electrode accessories
9. Installation & Commissioning
The Supplier shall provide:
Installation
Functional testing
Output testing
Channel testing
Safety testing
User demonstration
10. Training
The Supplier shall provide basic training covering:
Operation
Selection of stimulation mode
Intensity adjustment
Electrode placement
Treatment timer
Cleaning
Basic troubleshooting
Minimum training:
1 working day
11. Warranty
Minimum warranty:
24 months from commissioning
The warranty shall cover the stimulator, power supply, cables and supplied accessories.
12. Documentation
The Supplier shall provide:
User manual
Technical specifications
Safety certificate
Regulatory certificate
Warranty certificate